Saudi Arabia 2007/2008  Calibration Checks and Quality Assurance  By  David Delene Christopher Kruse
Saudi Arabia 2007/2008
Calibration Checks
and Quality Assurance
By
David Delene
Christopher Kruse

---
15 um Bead Checks
15 um Bead Checks

---
30 um Bead Checks
30 um Bead Checks

---
August 1, 2008 calibration check on the NCAR FSSP (SN 277-0676-06) using 15 µm beads. Calibration check was performed at UND before being shipped to Saudi Arabia.

---
September 3, 2008 calibration check on the NCAR FSSP (SN 277-0676-06) using 15 µm beads. Calibration check was preformed while the probe was in Saudi Arabia on the Research Aircraft.

---
September 3, 2008 calibration check on the WMI FSSP (SN 1947-0281-60) using 15 µm beads. Calibration check was preformed while the probe was in Saudi Arabia on the Research Aircraft.

---
August 15, 2008 calibration check on the WMI FSSP (SN 1947-0281-60) using 15 µm beads. Calibration check was preformed while the probe was in Saudi Arabia on the Research Aircraft.

---
FSSP Data
FSSP Data
Show use a vacuum to draw air through the FSSP for calibration checks.
Need to do a evaluation of the NCAR FSSP once it returns from Saudi Arabia.
Need to look closely at the 30 um bead checks.

---
Dew Point Temperature
Dew Point Temperature
Unedited data from the July 9, 2008 Saudi Arabia flight. The temperature (degrees C) is black, dew point temperature (degrees C) is green, and the FSSP concentration (#/cm^3) data is blue.

---
Dew Point Temperature
Edited data from the July 9, 2008 Saudi Arabia flight. The temperature (degrees C) is black, dew point temperature (degrees C) is green, and the FSSP concentration (#/cm^3) data is blue.

---
Temperature
Temperature
Unedited temperature data from the June 28, 2008 Saudi Arabia flight.

---
Temperature
Edited temperature data from the June 28, 2008 Saudi Arabia flight.

---
Winds
Winds
The north/south component of the horizontal winds from the AIMMS probe on August 8, 2008. The first and last intervals are when the aircraft is heading north, the two intervals in the middle are when the aircraft is heading south.

---
Winds
The east/west component of the horizontal winds from the AIMMS probe on August 8, 2008. The first and last intervals are when the aircraft is heading north, the two intervals in the middle are when the aircraft is heading south.

---
Winds
The magnitude of the horizontal winds from the AIMMS probe on August 8, 2008. The first and last intervals are when the aircraft is heading north, the two intervals in the middle are when the aircraft is heading south.

---
Winds
The direction of the horizontal winds from the AIMMS probe on August 8, 2008. The first and last intervals are when the aircraft is heading north, the two intervals in the middle are when the aircraft is heading south.

---
Conclusion
Conclusion
AIMMS horizontal winds are not with in 1 m/s accuracy.
Temperature, Dew Point Temperature, FSSP, and PCASP data have been reviewed and necessary edits applied to create a quality assured data set.
Need to confirm conclusion from FSSP calibration checks by laboratory checks of the NCAR FSSP and a DMT calibration of the WMI FSSP.

---
Future Work
Future Work
Compute summary statistics on the FSSP data (total concentration, median volume diameter) at cloud base and near cloud top (-5 to -10 C)
Computer summary statistics on the PCASP data (total concentration, sub-micrometer concentration) at cloud base.
Continue to evaluate the AIMMS data. Need equations and software. Need to record raw data.
Evaluate Uwyo CCN counter data.
Process and evaluate 2-DC data.

---
Saudi Arabia 2007/2008  Cloud Microphysics  By  David Delene
Saudi Arabia 2007/2008
Cloud Microphysics
By
David Delene

---
Saudi Arabia Winter 2008
Saudi Arabia Winter 2008
The 1 Hz cloud droplet concentration at standard temperature and pressure versus pressure altitude for all January-April 2008 Saudi Arabia flights measurements with concentrations above 50 # cm-3.

---
Mali, West Africa
Mali, West Africa
2006
2007
The 1 Hz cloud droplet concentration at standard temperature and pressure versus pressure altitude for all Mali 2006 (Top-Blue) and Mali 2007 (Bottom-Red) measurements with concentrations above 50 # cm-3.

---
FSSP Concentrations
FSSP Concentrations
The 1 Hz cloud droplet concentration at standard temperature and pressure for all measurements with concentrations above 50 # cm-3.
898
547
292
122
59
All
Saudi S08
651
240
93
65
53
All
Saudi W08
554
319
191
113
62
Top
Mali 2007
417
255
136
78
56
Top
Mali 2006
623
388
227
119
62
Bottom
Mali 2007
867
587
384
182
67
Bottom
Mali 2006
570
333
195
110
60
All
Mali 2007
841
505
277
125
61
All
Mali 2006
95%
75%
50%
25%
5%
Location
Date

---
Saudi Arabia Winter 2008
The 1 Hz liquid water from the DMT hot wire probe for all January-April 2008 Saudi Arabia flights measurements with FSSP total concentrations above 50 # cm-3.

---
Saudi Arabia Summer 2008
Saudi Arabia Summer 2008
The 1 Hz liquid water from the DMT hot wire probe for all June and July 2008 Saudi Arabia flights measurements with FSSP total concentrations above 50 # cm-3.

---
Hot Wire Liquid Water
Hot Wire Liquid Water
The 1 Hz liquid water from the DMT hot wire probe for all flights measurements with FSSP total concentrations above 50 # cm-3.
0.59
0.14
0.05
0.02
0.00
All
Saudi W08
0.63
0.25
0.10
0.00
-0.02
All
Saudi S08
3.20
1.80
0.81
0.19
0.00
All
Mali 2007
1.46
0.62
0.24
0.09
0.00
All
Mali 2006
95%
75%
50%
25%
5%
Location
Date

---
Conclusions
Conclusions
??

---
Saudi Arabia Winter 2008
The 1 Hz cloud droplet concentration at standard temperature and pressure versus pressure altitude for all January-April 2008 Saudi Arabia flights measurements with concentrations above 50 # cm-3.