`
Cloud Spectrum
Measurements
By David Delene

---
Forward Scattering Spectrometer Probe (FSSP) On the Left Wing
 of the King Air 200 Research Aircraft

---
Optical Path of the FSSP
The beam splitter divides the scattered light onto two photodetectors.
One photodector is optically masked to not receive scattered light from near the laser beam’s center of focus.
Droplets are rejected as being out of the depth of field when the signal from the masked detector exceeds that from the unmasked detector.
FSSP schematic is taken from Dye and Baumgarnder, [1984]

---
FSSP Effective Sample Volume
Sample Volume = TAS*DOF*BD*(Tc/Ts)‏
TAS – Aircraft True Air Speed (~100 m/s)‏
DOF – FSSP Depth of Field (~2.9 mm)‏
BD – Laser Beam Diameter (~0.2 mm)‏
Tc – Number of Droplets Sized
        (Total Counts)‏
Ts – Number of Droplets within the DOF
        (Total Strobes)‏
Air Flow
Laser Beam
DOF
BD

---
Laser Beam Fraction Correction
Effective Laser Beam Diameter (Tc/Ts)‏
Total Strobes (Ts)‏
Laser Beam
Total Counts (Tc)‏
The effective laser beam diameter is the fraction of the total diameter where droplets are within the laser beam long enough so they can be sized.
A running average of droplet transit time through the beam is maintained. If the droplet time within the laser beam is less than the average, it is rejected from sizing but included in the running average.

---
FSSP Velocity Acceptance Ratio
The velocity acceptance ratio is based on the ratio of total FSSP counts to total FSSP strobes. Dye and Baumgarnder [1984] state that the theoretical velocity acceptance ratio is 62%.
The velocity acceptance ratio is based on the ratio of total FSSP counts to total FSSP strobes. Dye and Baumgarnder [1984] state that the theoretical velocity acceptance ratio is 62%.

---
Coincidence and Deadtime Corrections
cf – Correction factor
Fa – Activity Fraction
Fa – Activity Fraction
The 0.73 constant is an empirical factor found from computer simulations which takes into account particles which are still in the beam at the end of a reset delay period. This factor is described by Baumgardner [1983] and Baumgardner et al [1985].

---
FSSP Particle Loss Correction
Percentage of particle losses based on the measured FSSP activity.
Percentage of particle losses based on the measured FSSP activity.

---
Theoretical Loss Correction
Plot of the expected loss of the NCAR Research Aviation Facility's FSSP probe flying at 100 m/s. Figure taken from the NCAR Research Aviation Facility Bulleting No. 24.

---
FSSP Mie Function
FSSP Mie Function is taken from Dye. and Baumgarnder, [1984]

---
FSSP Calibration Procedure
The FSSP is calibrated to determine the instruments depth of field, laser beam diameter, and channel size boundaries. The channel counts obtained from measurements on beads of known size are used to determine the FSSP channel boundaries.
[Speaker Notes: The setfssp.c program is used to convert the information on the PMS calibration sheet to the FSSp channel boundaries. The setfssp.c program takes into account aspects of our individual probe.]

---
FSSP Bead Calibration Check
January 29, 2008 FSSP calibration check at 8:40:28 using 15 µm beads.

---
Why Measure Liquid Water Content?
Determines the Potential of Enhancing Precipitation using Cloud Seed Techniques (Mali, Saudi Arabia)‏
Basic Cloud Parameter (MPACE)‏
Icing Studies (WISP04, Sikorsky)‏
Comparison with Remote Sensing	Measurements (THORpex, IOP1)‏

---
Liquid Water Content Calculation
The ¶amount of liquid water for a given volume of air may be determined through mass integration of the cloud droplet distribution.
size 12{ ital "LWC"= (¶ over 6 ) ρ rSub { size 8{w} } Sum cSub { size 8{i=1} } cSup { size 8{m}} {N rSub { size 8{i} }} d rSub { size 8{i} } rSup { size 8{3} } }
ρw – Density of Water
Ni – Concentration of droplets in size channel i
di – Droplet diameter in size channel i
m – Total number of channels

---
March 10, 2004 Citation Flight
Time [sfm-UTC]
Liquid Water Content [g/m3]
King Probe - 0.218 ± 0.096
FSSP - 0.203 ± 0.085
4 Hz averaged FSSP (Black line) and King Probe (Red Line) cloud liquid water content data.

---
September 24, 2004 Citation Flight
Time [sfm-UTC]
Liquid Water Content [g/m3]
FSSP – 0.077 ± 0.149
King Probe - 0.065 ± 0.123
4 Hz averaged FSSP (Black line) and King Probe (Red Line) cloud liquid water content data.

---
How well do they Compare?

---
How well do they Compare?

---
Ice Contamination
— FSSP
— King Probe
The FSSP LWC calculations assume spherical water droplets.

---
Rapid Visualization of Data
— FSSP
— King Probe

---
One Last Plot
— FSSP
— King Probe

---
Conclusions
Beam fraction, coincidence, and dead time corrections need to be applied to the FSSP data to obtain accurate LWC measurements; however, apparently no airspeed correction has to be applied to the FSSP data.
The FSSP liquid water content agrees with the King Probe LWC in ice free conditions. Cases from two different field programs found FSSP to King ratios of 0.96 and 1.20.
Well written software can automatically post process the FSSP data to provide accurate plots of the FSSP measurements during field projects.

---
Future Work (Assignment)‏
Compare the M300 real time processing of FSSP number concentration data with the post-processed FSSP number concentration data.
Use the January 10, 2008 flights for comparison.
Present results on March 5, 2008 at 1:00 p.m.
Mark it on the calendar, “FSSP data processing comparison”

---
     Any Questions?
[Speaker Notes: Lecture prepared by David Delene for UND physical met. class Spring 2003. Picture include aerosol balloon package launch in Feb. 1998. The University of Wyoming balloon-borne CCN Counter. The other images are from TSI, A 3010 CPC, 3025 ultrafine counter, and a 3080 Electrostatic Classifiers (www.tsi.com).]